## Cover Page
Capital Guide
Technology Report
Extracorporeal Perfusion Pumps and IABP Systems
Summary
Vendors
► Abbott | ► Fresenius | ► Medtronic
► ABIOMED | ► Getinge | ► Spectrum
► APMTD | ► LivaNova | ► Teleflex
Devices
Circulatory Assist Units, Cardiac, Intra-Aortic Balloon
Pumps, Extracorporeal Perfusion
Contents
Feature Comparison
Price Comparison
Member Interest
Vendor Contact Information
Hazards/Recalls
Prepared By
Mike Green
Phone: (800) 998-3274 ext. 5323
Email: mgreen@ECRI.org
Last Updated: May 28, 2024
In response to the FDA's recommendation to transition away from Getinge cardiovascular devices, many healthcare providers are seeking information on alternative extracorporeal perfusion pumps and intra-aortic balloon pump (IABP) systems. This report includes a brief comparison of the key features, average prices, and market interest for extracorporeal perfusion pumps used in extracorporeal membrane oxygenation (ECMO) circuits and for IABP systems.

## Features
Feature Comparison
Extracorporeal Perfusion Pumps | Abbott CentriMag | ABIOMED Breethe OXY-1 | APMTD ANIVIA SG1000 | Fresenius Novalung | Getinge CARDIOHELP (FDA recommends alternatives) | Getinge ROTAFLOW II | LivaNova LifeSPARC (discontinued by end of 2024) | Medtronic Bio-Console 560 | Medtronic VitalFlow | Spectrum Quantum Elite
FDA 510(k) Clearance Date | 2008 | 2020 |  | 2020 | 2011 | Not FDA-cleared for use in the United States | 2019 | 2005 |  | 2020
FDA 510(k) Cleared for ECMO | Yes (≤6 hrs) | Yes (≤6 hrs) | No | Yes (>6 hrs) | Yes (≤6 hrs) | No | Yes (≤6 hrs) | No | Yes (≤48 hrs) | No
FDA 510(k) Cleared for Cardiopulmonary Bypass | Yes (≤6 hrs) | Yes (≤6 hrs) | Yes (≤6 hrs) | Yes (>6 hrs) | Yes (≤6 hrs) |  | Yes (≤6 hrs) | Yes (≤6 hrs) | Yes (≤48 hrs) | Yes (≤6 hrs)
Indicated for Adults Only | No | No | No | Yes | No |  | Yes | No | Yes | No
Touchscreen Control | No | Yes | Yes | Yes | Yes | Touchscreen/rotary dial | No | No | Yes | Yes
Display Size, diagonal in | 12.1 for optional Mag Monitor | Not specified | 12.1 | 10.4 | 5.7 | Not specified | Not specified | 10.4 | Not specified | 12.1
Displayed Sensor Values | 1 blood flow, 3 pressures | 1 blood flow, 1 sweep gas flow, 1 air bubble | 1 blood flow, 1 air bubble, 2 pressures, 2 temperatures | 1 blood flow, 3 pressures, 1 air bubble | 1 blood flow, 4 external pressures, 1 air bubble, 1 level, 2 external temperatures, 3 internal pressures, 2 internal temperatures, 1 venous oxygen saturation (SvO2), 1 hemoglobin (Hb), 1 hematocrit (Hct) | 1 blood flow, 2 air bubbles | 1 blood flow | 1 blood flow, 1 air bubble, 2 levels | 1 blood flow, 1 air bubble | 2 blood flows (venous, arterial), 4 pressures, 4 temperatures, 1 venous oxygen saturation (SvO2), 1 arterial oxygen saturation (SaO2), 1 hemoglobin (Hb), 1 hematocrit (Hct)
Oxygen Source | External | Integral | External | External | External | External | External | External | External | External
HL7-Compliant Patient Monitoring | No | No | No | No | No | No | No | No | No | Optional
Pump Type | Centrifugal | Centrifugal | Centrifugal | Centrifugal | Centrifugal | Centrifugal | Centrifugal | Centrifugal | Centrifugal | Centrifugal
Pump Motor Location | Separate from controller | Separate from controller | Separate from controller | Separate from controller | Integrated with controller | Separate from controller | Separate from controller | Separate from controller | Separate from controller | Separate from controller
Max Speed, rpm | 5500 | 4500 | 5000 | 10000 | 5000 | 5000 | 7500 | 4000 | 3800 | Not specified
Max Flow Rate, lpm | 1.5 PediMag, 10 CentriMag | 5 | 10 | 7 | 5 or 7 | 7 | 8 | 10 | 7 | 8
Emergency Hand Crank | No | Yes | Yes | No | Yes | Yes | No | Yes | Yes | Yes
Air/Ground Transport | Yes | No | No | No | Yes | Not specified | No | No | No | No
Battery Type | Li-Ion | Li-Ion | LiFePO4 | Li-Ion | Li-Ion | Li-Ion | Li-Ion | Sealed lead-acid | Not specified | Li-Ion
Battery Life, hrs | 2.5 | 3.3333 | 1 | 2 | 1.5 | 1.5 | 2 | 0.5 | Not specified | 1.5
Battery Recharge Time, hrs | 5 | 6.5 | Not specified | Not specified | 5 | Not specified | 6 | 18 to 24 | Not specified | Not specified
Line Power, VAC | 100-240 | 100-240 | 85-264 | 100-240 | 100-240 | 100-240 | 120-240 | 100-240 | Not specified | Not specified
Controller Weight, lbs | 13 | 47 | 26.5 | 26 | 22 | 13.9 | 18.3 | 37.9 | Not specified | 15.5
Controller Dimensions, H x W x D, in | 13 x 10.5 x 3.9 | 26 x 14 x 9 | 13.8 x 9.4 x 4.3 | 12 x 11 x 15 | 17 x 9.8 x 11.8 | 8.5 x 9 x 17 | 12 x 8.5 x 8 | 12.5 x 8.9 x 43 | Not specified | Not specified
Warranty, yrs | 1 | 1 | 1 | 1 | 1 | Not specified | 1 | 1 | Not specified | Not specified
IABP Systems | Getinge CARDIOSAVE Hybrid (FDA recommends alternatives) | Teleflex AC3 Optimus
Touchscreen Control | Yes | Yes
Display Size, diagonal in | 12.1 | 13.3
Displayed Waveforms | ECG, AP, balloon pressure | ECG, AP, balloon pressure
Trigger Modes | ECG, AP, pacer A, pacer V/AV, internal | ECG pattern, ECG peak, AP, pacer A, pacer V, internal, Afib
Max Pump Rate, bpm | 200 | 200
Frequency Weaning | 1:1, 1:2, 1:3 | 1:1, 1:2, 1:4, 1:8
Pressure Range, psi | 2200 | 500 to 2,900
IAB Catheter Sizes, Fr | 7, 7.5, 8 | 7, 7.5, 8, 9
IAB Sizes, cc | 25, 30, 34, 40, 50 | 30, 35, 40, 50
In Vivo Calibration | Yes | No
Helium Cylinder Sizes, L | Disposable: 140; Refillable: 99 | Disposable: 33, 150; Refillable: 100, 105
Battery Type | Li-Ion (hot-swappable) | Sealed lead acid
Battery Life, hrs | 3 | 3 (with two batteries)
Battery Recharge Time, hrs | 5 | 8
Line Power, VAC | 100-240 | 85-265
Console Weight, lbs | 114 | 104
Console Dimensions, H x W x D, in | 44 x 22 x 27 | 33.3 x 13 x 24.5
Warranty, yrs | 1 | 1
Note on D13: Green, Michael: External oxygen source preferred (integral oxygen source intended as a backup)
Note on E7: Lee, Kevin: ECMO would be off-label use
Note on J7: Lee, Kevin: ECMO would be off-label use
Note on L7: Lee, Kevin: ECMO would be off-label use
Note on L14: Green, Michael: Requires Quantum Informatics LIVE VUE software license
Special Note
In May 2024, the FDA recommended that health care facilities transition away from use of the Getinge CARDIOHELP (with HLS Sets) and CARDIOSAVE Hybrid/Rescue systems. If possible, providers should use alternative devices.
These recommendations were based on safety concerns related to 12 recalls for the CARDIOSAVE systems and 8 recalls for the CARDIOHELP systems since 2023. According to the FDA, the manufacturer has not sufficiently addressed the problems associated with these devices.
For providers that do not have alternatives and continue to use the CARDIOHELP and CARDIOSAVE systems, the FDA recommends reviewing all urgent medical device correction notices from Getinge.
Furthermore, Getinge announced that US sales for both systems will be limited to customers that do not have alternatives. The manufacturer also plans to file for FDA clearance for the next generation CARDIOHELP 2 and CARDIOSAVE Plus systems by mid-2025.
Additional Vendors
There are no additional US vendors that offer extracorporeal perfusion pumps and IABP systems.
Also, note that while clinicians may choose to use a competitor's IABs to accommodate patients of different sizes, it is preferable to use balloons from the same manufacturer as the IABP console.

## Pricing
Price Comparison
HIDE THESE COLUMNS
| Device 1 | Abbott CentriMag | APMTD ANIVIA SG1000 | Fresenius Novalung | Getinge CARDIOHELP | Getinge ROTAFLOW | LivaNova LifeSPARC | Medtronic Bio-Console 560 | Medtronic VitalFlow | Spectrum Quantum Elite | Vendor 10 Model 10
ECRI Pricing Data
| Average List Price | 52258 |  |  | 142470 |  | 50000 | 43100 |  | 199920
| Average Quoted Price | 44545 | 39900 |  | 118911 |  | 25000 | 31525 | 85950 | 199920
| Device 2 | Getinge CARDIOSAVE Hybrid | Teleflex AC3 Optimus
ECRI Pricing Data
| Average List Price | 95575 | 98229
| Average Quoted Price | 62202 | 58226
Prices for the extracorporeal perfusion pumps reflect the pump speed controller, motor, and accessories.
The Spectrum Quantum Elite pricing also includes an electronic gas blender and blood-gas analyzer which has an average quoted price of $64,750.
No recent pricing information is available for the ABIOMED Breethe OXY-1 and Fresenius Novalung. Based upon data from 2021 to 2022, the estimated cost for these alternatives would be $50,000 to $60,000.
Prices for the IABP systems includes the console and accessories.

## Supply Guide Data (hidden)
Supply Guide Data
Vendor | Item | Part # | Oxygenator Included | Cannulas Included | UOM | Qty/UOM | Low | 25th Percentile | Median | 75th Percentile | High | Member(s) Submitting | National Spend
Abbott | CentriMag Blood Pump | 102953 | No | No | EA | 1 | 8500 | 10500 | 12300 | 12600 | 13000 | 108 | 16039040
Abbott | PediMag Blood Pump | 20190052 | No | No | EA | 1 | 6500 | 8000 | 10000 | 10000 | 10000 | 22 | 1660500
ABIOMED | OXY-1 Pump Lung Unit (PLU) | 0010400100 | Yes | No | EA | 1 | No recent data | No recent data | No recent data | No recent data | No recent data | 0 | 0
Fresenius | XLung Kit | F32000004 | Yes | No | EA | 1 | 7921 | 7921 | 7921 | 8900 | 8900 | 4 | 723036
Getinge | HLS Set Advanced 7.0 | BEQHLS7050 | Yes | No | EA | 1 | 9017.65 | 11500 | 12552 | 13860 | 14500 | 181 | 2122664
Getinge | HLS Set Advanced 5.0 | BEQHLS5050 | Yes | No | EA | 1 | 9500 | 11500 | 11800 | 13860 | 14500 | 31 | 1310836
Getinge | PLS Set | 70106.8020 | Yes | No | EA | 1 | No recent data | No recent data | No recent data | No recent data | No recent data | 0 | 0
LivaNova | TandemLife LS | 58402415 | Yes | Yes | EA | 1 | 16500 | 16500 | 16500 | 16500 | 16500 | 4 | 280500
LivaNova | TandemLife LS | 58402417 | Yes | Yes | EA | 1 | 15500 | 16500 | 16500 | 18000 | 19500 | 23 | 1451000
LivaNova | TandemLung LS | 58302916 | Yes | Yes | EA | 1 | 21500 | 21500 | 21500 | 24000 | 24000 | 4 | 179500
LivaNova | TandemLung LS | 58303118 | Yes | Yes | EA | 1 | 21500 | 21500 | 24000 | 24000 | 24000 | 6 | 404500
Medtronic | Affinity CP | AP40 | No | No | CS | 4 | 690.12 | 690.12 | 690.12 | 690.12 | 690.12 | 1 | 37266.48
Medtronic | Affinity CP with Balance Biosurface | BBAP40 | No | No | CS | 4 | 642.6 | 642.6 | 679.8 | 803.4 | 1212 | 18 | 80998.48
Medtronic | Affinity CP with Cortiva BioActive Surface | CBAP40 | No | No | CS | 4 | 684 | 867.8 | 867.8 | 982.24 | 1153.6 | 19 | 94128.96
Spectrum | Qura CP22 | CP22VVT | No | No | EA | 1 | 3813 | 4013 | 4107 | 4200 | 4500 | 12 | 2529750
This shape represents a table slicer. Table slicers are supported in Excel or later.
If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.
This shape represents a table slicer. Table slicers are supported in Excel or later.
If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.
The prices below were recorded in the last 13 months by our Supply Guide service for the centrifugal pump set, comprised of the disposable pump head and in some cases other extracorporeal perfusion supplies (e.g., oxygenator, cannulas).
This shape represents a table slicer. Table slicers are not supported in this version of Excel.
If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.

## Interest
Member Interest
Previous version
| Vendor | Quotes | %
| Abbott CentriMag | 16 | 33.33%
| Getinge CARDIOHELP | 14 | 29.17%
| Spectrum Quantum Elite | 7 | 14.58%
| Medtronic Bio-Console 560 | 6 | 12.5%
| LivaNova LifeSPARC | 5 | 10.42%
| ABIOMED Breethe OXY-1 | 0 | 0%
| Fresenius Novalung | 0 | 0%
| n=48 | 48 | 100%
| Model | Quotes | %
| Getinge CARDIOHELP | 15 | 30%
| Abbott CentriMag | 13 | 26%
| Spectrum Quantum Elite | 13 | 26%
| LivaNova LifeSPARC | 5 | 10%
| Medtronic Bio-Console 560 | 2 | 4%
| APMTD ANIVIA SG1000 | 1 | 2%
| Medtronic VitalFlow | 1 | 2%
| ABIOMED Breethe OXY-1 | 0 | 0%
| Fresenius Novalung | 0 | 0%
| n=50 | 50 | 100%
| Model | Quotes | %
| Teleflex AC3 Optimus | 38 | 52.78%
| Getinge CARDIOSAVE Hybrid | 34 | 47.22%
| n=72 | 72 | 100%
Market interest charts are based solely upon quotations submitted to ECRI by members of the Capital Guide advisory service. This data is not validated market share and is time sensitive.
Member interest in extracorporeal perfusion pumps excludes multi-pump heart-lung machines.

## Vendors
Vendor Contact Information
Abbott Laboratories Inc
100 Abbott Park Rd | Phone: | (224) 667-6100
Abbott Park, IL 60064
United States
http://www.abbott.com
Models:
CentriMag
ABIOMED Inc
22 Cherry Hill Dr | Phone: | (978) 646-1400
Danvers, MA 01923 |  | (800) 422-8666
United States
http://www.abiomed.com
Models:
Breethe OXY-1
APMTD Inc
226 Lowell St Unit B16 | Phone: | (617) 275-4019
Wilmington, MA 01887
United States
http://www.apmtd.com
Models:
ANIVIA SG1000
Fresenius Medical Care North America
920 Winter St | Phone: | (781) 699-9000
Waltham, MA 02451-1457 |  | (800) 662-1237
United States
http://www.fmcna.com
Models:
Novalung
Getinge USA Inc
1 Geoffrey Way | Phone: | (888) 943-8872
Wayne, NJ 07470
United States
http://www.getinge.com
Models:
CARDIOHELP
CARDIOSAVE Hybrid
LivaNova plc
20 Eastbourne Terrace | Phone: | 44 (203) 7865275
London, W2 6LG
England
http://www.livanova.com
Models:
LifeSPARC
Medtronic Cardiac Rhythm and Heart Failure Management Div Medtronic Inc
8200 Coral Sea St NE | Phone: | (800) 722-3278
Mounds View, MN 55112
United States
http://www.medtronic.com
Models:
Bio-Console 560
VitalFlow
Spectrum Medical Inc
481 Munn Rd Suite 180 | Phone: | (800) 265-2331
Fort Mill, SC 29715
United States
http://www.spectrummedical.com
Models:
Quantum Elite
Teleflex Medical Arrow International Inc
16 Elizabeth Dr | Phone: | (978) 250-5100
Chelmsford, MA 01824 |  | (800) 343-3297
United States
https://www.teleflex.com/usa/en/
Models:
AC3 Optimus

## Hazards
Hazards/Recalls
Search date range: May 2023 - May 2024
Model | Critical Priority Alert | High Priority Alert | Normal Priority Alert
Extracorporeal Perfusion Pumps
Abbott CentriMag | 0 | 1 | 0
ABIOMED Breethe OXY-1 | 0 | 0 | 0
APMTD ANIVIA SG1000 | 0 | 0 | 0
Fresenius Novalung | 0 | 0 | 0
Getinge CARDIOHELP | 0 | 8 | 1
LivaNova LifeSPARC | 0 | 0 | 1
Medtronic Bio-Console 560 | 0 | 0 | 0
Medtronic VitalFlow | 0 | 1 | 0
Spectrum Quantum Elite | 0 | 0 | 0
IABP Systems
Getinge CARDIOSAVE Hybrid | 3 | 3 | 0
Teleflex AC3 Optimus | 2 | 0 | 0
Model | Accession # | Published | ECRI Priority | Alert
AC3 Optimus | A42358 | 2024-05-03 | Critical | Teleflex—Arrow FiberOptix and UltraFlex Intra-Aortic Balloon Catheter Kits: Balloon May Fail to Inflate, Central Lumen Segment Contained in Balloon May Be Damaged or Broken, and/or Helium Loss or Blood in the Helium Pathway
AC3 Optimus | A39841 02 | 2023-05-05 | Critical | Teleflex—Arrow AutoCAT 2 and AC3 Optimus Intra-Aortic Balloon Pumps: May Exhibit Short Battery Run-Times and Loss of Power during Use [FDA Class I]
CARDIOHELP | A42454 | 2024-05-23 | High | Maquet/Getinge—Cardiohelp-i Systems: Measured Patient Leakage Current May Exceed Expected Levels
CARDIOHELP | H0931 | 2024-05-20 | High | FDA Recommends Transitioning Away from Use of Getinge/Maquet Cardiosave Hybrid IABPs, Cardiosave Rescue IABPs, Cardiohelp Systems, and HLS Sets
CARDIOHELP | A41779 02 | 2024-02-01 | High | Maquet/Getinge—HLS Sets: Sterility of Emergency Priming Line May Be Compromised​
CARDIOHELP | A41552 | 2023-11-22 | High | Maquet/Getinge​—Cardiohelp-i Systems: Venous Probe Connection Cable May Break
CARDIOHELP | A41515 | 2023-11-13 | High | Maquet/Getinge​—Cardiohelp-i Systems: Manufacturer Updates Instructions for Use
CARDIOHELP | A41365 02 | 2023-10-27 | High | Maquet​/Getinge—Cardiohelp Emergency Drives: Rotation of Hand Crank May Be Blocked or Impaired
CARDIOHELP | A41237 02 | 2023-10-27 | High | Maquet/Getinge​—Cardiohelp-i Systems: Nut Bolt May Break off of Slide Rail Plate
CARDIOHELP | A40107 04 | 2023-07-13 | High | Maquet/Getinge—HLS Sets: Sterility May Be Compromised
CARDIOHELP | A41931 | 2024-02-09 | Normal | Maquet/Getinge​—Bubble Sensors Used with Cardiohelp-i Systems: Label May Contain Incorrect UDI
CARDIOSAVE Hybrid | H0931 | 2024-05-20 | High | FDA Recommends Transitioning Away from Use of Getinge/Maquet Cardiosave Hybrid IABPs, Cardiosave Rescue IABPs, Cardiohelp Systems, and HLS Sets
CARDIOSAVE Hybrid | A40044 02 | 2023-12-05 | Critical | Datascope/Getinge—Cardiosave Hybrid and Cardiosave Rescue Intra-Aortic Balloon Pumps: Blood Ingress May Cause Unexpected Shutdown
CARDIOSAVE Hybrid | A41072 | 2023-08-08 | Critical | Datascope/Getinge—Cardiosave Hybrid and Cardiosave Rescue Intra-Aortic Balloon Pumps: May Exhibit Various Problems
CARDIOSAVE Hybrid | A40805 | 2023-06-07 | Critical | Datascope/Getinge—Cardiosave Hybrid and Cardiosave Rescue Intra-Aortic Balloon Pumps: Blood within Perforated IAB Catheter May Enter IABP Console
CARDIOSAVE Hybrid | A41071 | 2023-08-08 | High | Datascope/Getinge—Cardiosave Hybrid and Cardiosave Rescue Intra-Aortic Balloon Pumps: May Fail to Dock and Charge Battery; May Exhibit Poor or No ECG Signal
CARDIOSAVE Hybrid | A40835 | 2023-06-13 | High | Datascope/Getinge—Cardiosave Hybrid and Cardiosave Rescue Intra-Aortic Balloon Pumps: May Lose Ability to Charge Batteries; Unexpected Shutdown May Occur
CentriMag | H0913 | 2024-02-26 | High | Abbott—Thoratec CentriMag Acute Circulatory Support Systems: Pump Motor Can Spontaneously Become Hot Enough to Damage Blood
LifeSPARC | A42016 | 2024-03-01 | Normal | LivaNova—LifeSPARC and Hemolung Platforms: Manufacturer Discontinues Product​s
VitalFlow | A41457 | 2023-11-03 | High | Medtronic—Nautilus ECMO Oxygenators and Nautilus Smart ECMO Modules with Balance Biosurface: May Exhibit Degraded Heat Exchange Performance during Use

## Disclaimer
Policy Statement and Disclaimer
The information provided in this report is derived from several sources, including manufacturers, clinical users, and your facility. We make every reasonable effort to ensure that it is up-to-date and verify it with the most knowledgeable sources. This information has not been amplified, qualified, or confirmed by us through on-site observation, inspections, or review of hospital policies and procedures. The basis for our equipment recommendations, therefore, may not reflect all of the factors and perspectives provided through on-site, firsthand consultation and contact with all appropriate staff and departments.
The contents of this document should not be regarded as a guarantee of the safety, performance, or cost-effectiveness of any equipment, device, or system considered under this program. This analysis was created for your institution and is copyrighted. It may not be shared with manufacturers or any third party without written permission from ECRI.
©2024 ECRI. All Rights Reserved. This material is confidential, and use of the copyright notice is not intended to suggest or imply that the material has been published or is available for public dissemination. This report was created for your institution and should not be reproduced or transmitted to third parties without written permission from ECRI.